## Slide 1: First demonstration of portable Compton camera to visualize 223-Ra concentration for radionuclide therapy
Kazuya Fujieda (Waseda University)
J. Kataoka, S. Mochizuki, L. Tagawa, S. Sato, R. Tanaka (Waseda University),
K. Matsunaga, T. Kamiya, T. Watabe, E. Shimosegawa, J. Hatazawa(Osaka University),
S. Ohsuka (Hamamatsu Photonics K. K.)

## Slide 2: Outline
Introduction
- Radionuclide therapy
Experimental Methods
- Compton camera
Compton Camera Images of 223-Ra
- A small bottle
- NEMA IEC body phantom
Conclusion
1

## Slide 3: Introduction
Radiation therapies
RNT with alpha-emitting nuclides can minimize damage to normal cells
Hard to determine whether the implant has been properly delivered
2
External radiation therapies
Radionuclide therapy (RNT)
Irradiating beams from the outside of the body
Damage to
not only cancer cells
but also normal cells
Placing the radioactive implant in or near the tumor
Selectively damage cancer cells
Beta particles
- more than a few mm range
Alpha particles
- large linear energy transfer
- shorter range (~10 mm)

## Slide 4: 223-Ra
The characteristic of 223-Ra
Determine the position of 223-Ra by gamma rays emitted
3
An alpha-emitting nuclide
Used for treating bone metastasis cancer
Also emit gamma rays during its collapse process
Detect with a gamma-ray visualization module
Determine the position of 223-Ra
The main energy of gamma rays emitted from 223-Ra
- 154 keV, 271 keV, 270 keV,
351 keV, 402 keV

## Slide 5: SPECT vs. A Compton camera
Gamma-ray visualization modules in RNT
Reduce burdens to patients by using a Compton camera
4
Single photon emission tomography (SPECT)
Identify the direction of arrival of gamma rays by passive collimators
Burdens to patients
- The relatively long
exposure
- Planar geometry
A Compton camera
(A new method)
Identify the direction of arrival of gamma rays with Compton-scattering kinematics
Advantages to patients
- Wider field of view
- Less burdens
in imaging
- easiness of 3D imaging

## Slide 6: A Compton camera
The principle of a Compton camera
A Compton camera visualizes the arrival direction of gamma rays
5
Visualize the arrival direction of gamma rays () by utilizing Compton-scattering kinematics
is calculated from …
- The reaction position: ,
- The energy deposition:
- The formula of Compton scattering
Specify the position of source by overlapping a Compton cone
(whose vertical angle is )

## Slide 7: A Compton camera (DOI-CC)
DOI-CC
Imaging 223-Ra with a portable Compton camera (DOI-CC)
6
A portable Compton camera (16 cm × 15 cm × 15 cm, 2.5 kg)
Developed for an environmental survey in Fukushima
A Compton camera that can obtain the depth of interaction (DOI-CC)
Resolutions of DOI-CC at 662keV
- The angular resolution: ~8 (FWHM at the center of FOV)
- The energy resolution: 7.8 % (FWHM)
Constitution of DOI-CC
- Scintillator: Ce:GAGG arrays (Ce-doped Gd3Al2Ga3O12)
- Photon detector: 8×8 MPPC arrays
(Multi-pixel photon counter)
- Two couples of Ce:GAGG and MPPC @ scatterer
- One Ce:GAGG sandwiched with two MPPC @ absorber
Kishimoto et al., IEEE Trans. Nucl. Sci, 59, (2013), 38-43.
 | Scatterer | Absorber
Crystal size | 2.0 × 2.0 × 4.0 mm3 | 2.0 × 2.0 × 2.0 mm3
Array | 11 × 11 arrays 2 × 2 set | 11 × 11 arrays 2 × 2 set
Layer | 2 | 10
Configuration of DOI-CC

## Slide 8: Imaging (a small bottle filled with 223-Ra)
Experimental setup
Confirm whether DOI-CC can reconstruct the true position of 223-Ra
7
10cm
or
15cm
A small bottle filled with 223-Ra (0.56 MBq)
DOI-CC
Set the bottle at 4 positions:
10 cm 0 deg. / 10 cm 45 deg. /
15 cm 0 deg. / 15 cm 45 deg.
The time for each measurement: 20 min
Energy spectrum (all events)
324-351 keV
269 & 271 keV
84 keV
154 keV
Set two energy bands for image reconstruction:
260-310 keV (target 269 & 271 keV)
288-388 keV (target 324-351 keV)

## Slide 9: Imaging (a small bottle filled with 223-Ra)
Reconstruct the true position of 223-Ra with DOI-CC successfully
8
Reconstructed images
- MLEM (maximum likelihood expectation maximization) reconstruction images
324-351
keV
269 & 271
keV
10 cm 0 deg.
10 cm 45 deg.
15 cm 0 deg.
15 cm 45 deg.

## Slide 10: Imaging (NEMA IEC body phantom)
Reconstruct the 3D image of a complex phantom with DOI-CC
9
Experiment setup
- NEMA IEC body phantom (consisted of six spheres)
- Three of them were filled with 223-Ra (9 kBq/ml)
- Rotate DOI-CC around the phantom for 3D imaging
(at 8 positions)
- The time for each measurement: 30 min
The diameters of the spheres | The radioactivity of 223-Ra | The distance from the center of the phantom
13 mm （a small sphere） | 10.8 kBq | 50 mm
22 mm （a middle sphere） | 50.4 kBq | 55 mm
37 mm （a large sphere） | 238.5 kBq | 60 mm

## Slide 11: Imaging (NEMA IEC body phantom)
Reconstruct the 3D image at each energy bands
10
Energy spectrum (only coincidence events)
- Set the four energy bands:
269 & 271 keV
324-351 keV
401 & 404 keV
832 keV
The energy of gamma ray targeted | The energy bands set
269 & 271 keV | 248–308 keV
324-351 keV | 318-378 keV
401 & 404 keV | 380-429 keV
832 keV | 742-842 keV

## Slide 12: Imaging (NEMA IEC body phantom)
The spheres are reconstructed at the correct position
11
3D reconstructed images (324-351 keV, 12590 events)
Simple back projection (SBP)
MLEM (20 iterations)
MLEM (20 iterations, normalized)
Normalized to enhance low-intensity spheres

## Slide 13: Imaging (NEMA IEC body phantom)
The spheres are reconstructed at the correct position
11
3D reconstructed images (324-351 keV, 12590 events)
MLEM (20 iterations)
MLEM (20 iterations, normalized)
X [mm]
Y [mm]
Z [mm]
X [mm]
Y [mm]
Z [mm]
X [mm]
Y [mm]
Z [mm]
Normalized to enhance low-intensity spheres
Simple back projection (SBP)

## Slide 14: Imaging (NEMA IEC body phantom)
The best result is the image of 324-351 keV
12
3D reconstructed images of the other energy bands (20 MLEM iterations)
269 & 271 keV
(12321 events)
401 & 404 keV
(5046 events)
832 keV
(525 events)

## Slide 15: Imaging (NEMA IEC body phantom)
The best result was the image of 324-351 keV
12
3D reconstructed images of the other energy bands (20 MLEM iterations)
269 & 271 keV
(12321 events)
401 & 404 keV
(5046 events)
832 keV
(525 events)
X [mm]
Y [mm]
Z [mm]
X [mm]
Y [mm]
Z [mm]
X [mm]
Y [mm]
Z [mm]
Less number of events than 324-351 keV
Including fake events
(chance coincidence, etc.)
Less number of events than 324-351 keV

## Slide 16: Analysis (NEMA IEC body phantom)
Confirm the relative intensity of the spheres
13
Analyze the SBP reconstructed image of 324-351 keV
- Obtain an azimuth profile of a thin slice
(every 15 mm in the radial direction of the phantom)
- If there are three spheres in the slice,
there are three peaks in the profile
- Fit with three Gaussian functions and a constant offset
- (The integral value of each Gaussian function)
= (The relative radioactivity of each sphere)
Z
Y
X
CC
223
Ra
X [mm]
Y [mm]
Z [mm]

## Slide 17: Analysis (NEMA IEC body phantom)
Fit with three Gaussian function and constant offset
14
The azimuth profiles of the 324-351 keV SBP image
=15-30 mm (no sphere)
=45-60 mm (3 spheres)
=120-135 mm (no sphere)
Three peaks: Large (red),
Middle (green), Small (blue)
Only one peak
More than three peaks

## Slide 18: Analysis (NEMA IEC body phantom)
Reconstruct the relative intensity of 223-Ra successfully
15
The relative intensity of the spheres
- Compare with the integral value of the spheres where the spheres exist
The diameter of the spheres | The integral value | The normalized integral value (the value of the large sphere equal to the radioactivity of 223-Ra) | The radioactivity of 223-Ra
13 mm (small) | 101.2 +/- 24.5 | 13.0 +/- 3.1 | 10.8 kBq
22 mm (middle) | 381.3 +/- 65.0 | 48.9 +/- 8.3 | 50.4 kBq
37 mm (large) | 1858.4 +/- 142.4 | 238.5 +/- 18.3 | 238.5 kBq
Almost consistent with the actual 223-Ra intensity of each sphere

## Slide 19: Conclusion
The imaging of 223-Ra with DOI-CC was succeeded
16
A small bottle filled with 223-Ra
- Reconstruct the correct position of 223-Ra with DOI-CC successfully
A complex phantom of 223-Ra
- Obtain 3D reconstructed images
- Reconstruct the correct positon and relative intensity successfully

## Slide 20: Future Work
Image the body of a patient
Optimize the detector configuration to improve the angular resolution
17
DOI-CC
The experimental setup
10-min measurement
with DOI-CC
Compare with
the image of SPECT
(30-min measurement)
The reconstructed image
feet
side
head
side
[mm]
[mm]
The accumulation consistent with SPECT
Shorter time for measurement
Wider FOV
Preliminary

## Slide 22: Appendix (Analyze NEMA IEC body phantom)
The mean value is similar to the true value
A1
The mean value of Gaussian functions
(= the angle of the center of the spheres)
The area where three spheres exist is denoted by solid lines
Affected by the largest sphere
The diameters of the spheres | The angle of the center of the spheres
13 mm （A small sphere） | 450
22 mm （A middle sphere） | 210
37 mm （A large sphere） | 330
The diameters of the spheres | The angle of the center of the spheres
13 mm （A small sphere） | 
22 mm （A middle sphere） | 
37 mm （A large sphere） |