## Slide 1: Evaluation of a novel photon-counting CT system using 16-ch MPPC array for multicolor 3D imaging
Takuya Maruhashi (Waseda Univ.)
M. Arimoto (Kanazawa Univ.), J. Kataoka, H. Kiji (Waseda Univ.),
S. Terazawa, H. Nitta, S. Shiota (Hitachi Metal Ltd.), H. Ikeda (ISAS)
Maruhashi et al NIM-A, in press

## Slide 2: Outline
Introduction
Approach
3D multicolor imaging
K-edge imaging of a contrast agent
Summary & Future work
- PC-CT system using MPPC (Multi pixel photon counter)
- Identification of a lighter materials using a X-ray spectral information
- Reconstruction of a Iodine and gadolinium contrast agents

## Slide 3: Introduction
CT image
Image reconstruction
X-ray generator
Detector
Phantom
X-ray
GOS scintillator
photodiode (PD)
Conventional
CT detector
High radiation dose (~10mSv/CT scan)
- All X-ray signal piles up
Beam hardening artifact
- Miss identification of material
X-ray CT imaging
1
X-ray
Problems of conventional CT

## Slide 4: Photon counting CT (PC-CT)
Eth0
Eth2
Eth3
Eth1
Multiple energy thresholds
⇒ Enable to detect
individual X-ray pulses
Integrate X-ray signals
with a certain time
⇒ Readout current (DC) signal
Readout DC signal
Count each pulse signal
Conventional CT cannot acquire
the energy information
of individual X-ray pulses
Cut dark noise
⇒ Achieve low dose
2
Photon counting CT
Time
Charge
X-ray signals
Conventional CT system
Phantom
X-ray generator
Detector
X-ray

## Slide 5: Energy resolved CT image
PC-CT system can reconstruct CT images
in different energy bands with one CT scan
- Improve a contrast of soft tissue
- Reduce a beam hardening of X-ray
Multiple energy thresholds
- Energy discrimination of X-ray signals
Energy resolved CT image
20 – 40 keV
90 – 120 keV
60 – 90 keV
40 – 60 keV
PC-CT enables more accurate
diagnosis than conventional CT
3
Eth 1
Eth 2
Eth 3
Eth 4
4 Energy resolved CT images
are reconstructed
Ex) 4 energy thresholds

## Slide 6: MPPC PC-CT system
X-ray
Scintillator
Photodiode
High speed
scintillator
MPPC
Conventional CT
Photon counting CT
High avalanche gain (M~)
Fast time response (~10 ns/pulse)
We proposed a PC-CT system using MPPC (Multipixel photon counter)
Our PC-CT system using 16-channel MPPC array
LSI
X-ray
generator
…
…
16-ch MPPC & YGAG array
YGAG scintillator
~70 ns
4
MPPC performance

## Slide 7: Detailed setup of our PC-CT system
Scintillator
MPPC
Phantom
Cu Collimator
X-ray generator
Al filter (1mm)
LSI
16-ch signals
Our PC-CT system reconstruct 16 slice CT image
5
Ch 1
Ch 2
Ch 3
Ch 4
~
Ch 15
Ch 16
~
~
~
The projection data in each direction is acquired
by repeating the linear and rotation move
Arimoto et al NIM-A,
in press

## Slide 8: Schematic view of the signal processing in LSI
…
Eth0
Eth2
Eth3
Eth1
Ch 2
Eth0
Eth2
Eth3
Eth1
Ch 1
Newly developed LSI set 4 energy thresholds at each channel
Ch 1
Ch 2
Ch 3
Ch 4
~
Ch 15
Ch 16
~
~
LSI
6
16 ch
Schematic view of the signal processing
Energy discrimination in 16-ch
The signals of MPPC array

## Slide 9: Color imaging of a lighter phantom
X-ray
Monochromatic
Monochromatic image
No information about material
Conventional energy integrated image
Lighter phantom
Conventional CT cannot acquire the X-ray spectrum information
Conventional CT
Photon counting CT
7
Multicolor image
Accurate identification of material
Notes: 16ch 撮影を3度繰り返し、スライス画像を取得の説明

## Slide 10: Energy resolved image of a lighter phantom
X-ray
( Tube voltage : 150 kV , Tube current : 0.1 mA )
Eth 0,1,2,3 : 20, 30, 60, 90 keV
Slice energy resolved CT images
…
…
…
…
20 - 30 keV
30 - 60 keV
60 - 90 keV
90 - 150 keV
Using 4 energies resolved image,
we try to emphasize each lighter material
Plastic , Aluminum , Iron
Main parts of lighter
8

## Slide 11: Reconstructing multicolor image
20-30 keV
30-60 keV
60-90 keV
90-150 keV
Energy subtraction
Energy subtraction
Energy
resolved
…
…
…
…
…
…
…
Low energy
plastic
Medium energy
aluminum
High energy
iron
9
Subtracting energy resolved image
between 2 energy windows,
we reconstructed specific material
Slice energy resolved CT images
Using different energy dependence
of CT values, our PC-CT system decomposed the lighter materials.
CT values change considerably
in 2 energy windows
Energy subtraction

## Slide 12: 3D color lighter imaging
Aluminum
Iron
Plastic
…
…
…
keV
Intensity
10
Using multi slice images of lighter material,
our PC-CT system reconstructed
3D color image of
plastic, aluminum and iron parts
Plastic
Aluminum
Iron

## Slide 13: Overlay 3D lighter imaging
Overlay slice
CT images
x
y
z
XY
YZ
Aluminum
Iron
Plastic
11
RGB image
Projection of 3D image
RGB 3D image
Reconstruct
3D image

## Slide 14: K-edge imaging of contrast agent
Photon energy [keV]
Linear attenuation coefficient
Using the K-absorption edge of a contrast agents
Greatly improve the blood tissue contrast
PC-CT has great advantages in K-edge imaging
Simultaneous imaging of multiple contrast agents
Absolute quantification of contrast agent
- Iodine, Gd, AuNP …
- Concentration (mg/ml) of contrast agent
In this study, we evaluated our PC-CT system’s performance of K-edge imaging
12
K-edge imaging
PC-CT system achieves more detailed tissue diagnosis

## Slide 15: K-edge imaging of iodine contrast agent
Water
Iodine
15 mg/ml
35-45keV
45-55keV
(a)
(b)
(c) = (a) – (b)
keV
Linear attenuation coefficient
- Tube voltage : 120 kV
- Tube current : 0.1 mA
- Eth 0,1,2,3 : 35, 45, 55, 70 keV
Energy subtraction
K-absorption edge
Iodine : 33.2 keV
Using the K-absorption edge of iodine,
our PC-CT system improved the iodine contrast
Iodine concentration
5, 10, 15, 20 mg/ml
13
20 mg/ml
5 mg/ml
10 mg/ml

## Slide 16: Reconstruction of 16-ch density image
θ
…
Energy subtraction image
…
14
Estimate density ( ) of material
least-squares method
Projection of the density image
0
180
20 mg/ml
15 mg/ml
10 mg/ml
5 mg/ml
Density image
Degree
Density [mg/ml]

## Slide 17: Measured density vs Real density
Iodine density image
To measure the density value and error,
we defined region of interest (ROI)
5 mg/ml
10 mg/ml
15 mg/ml
20 mg/ml
…
Our PC-CT accurately reconstructed iodine phantom (5, 10, 15, 20 mg/ml)
15
ROI in each iodine phantom
Measured density [mg/ml]
Real density [mg/ml]
Ideal line
Plot : measured density

## Slide 18: Simultaneous imaging of iodine & gadolinium
water
Iodine 10mg/ml
Eth 0,1,2,3 : 35, 45, 55, 70keV
K-absorption edge : Iodine 33.2 keV , Gd 50.2 keV
70-120keV
55-70keV
45-55keV
35-45keV
Chart: 
| 10 | 15 | 20 | 30 | 40 | 50 | 50.24 | 50.24 | 60 | 80 | 100 | 150
Gd | 2125.1 | 734.7 | 339.7 | 110.6 | 54.668 | 30.494 | 30.099 | 146.94 | 92.43 | 44.003 | 24.49 | 8.69
Iodine | 801.618 | 271.6923 | 125.3699 | 42.2057 | 32.3063 | 176.6419 | 108.953 | 60.7376 | 37.3546 | 17.3043 | 9.5741
Difference of K-absorption edge & Energy resolved CT image
We tried simultaneous imaging of multiple contrast agents
Multiple contrast agents with different pharmacokinetic
probe different diseases in one CT scan
Gd
10mg/ml
keV
Linear attenuation coefficient
Energy resolved CT images
16

## Slide 19: Result of simultaneous imaging
17
70-120keV
55-70keV
45-55keV
35-45keV
Iodine
density image
Gadolinium
density image
12 ± 1 mg/ml
10 ± 1 mg/ml
…
4 Energy resolved CT images
Our PC-CT system accurately identified the iodine and gadolinium phantom
Enable to probe different pharmacokinetic of contrast agents
Overlay density image
Energy subtraction
at K-absorption edge
of iodine and gadolinium
Iodine
Gd
Energy subtraction + least-squares method

## Slide 20: Summary & Future work
Summary
Future work
Our PC-CT system reconstructed 3D multicolor CT image of the lighter phantom
- Using the X-ray spectral information, we accomplished accurate material
identification
Our PC-CT system enables K-edge imaging of contrast agent
- Our system accurately reconstructed the concentration of Iodine and
gadolinium phantom
To more practical CT imaging, the number of channel : 16-ch → 64-ch
CT scan using a medical phantom or living mouse
18

## Slide 21
appendix

## Slide 22: PD, APD vs MPPC
Comparison of CT images (PD, APD and MPPC)
Check improvement of S/N by internal amplification
PD,APD
Size
1mm×1mm
Current readout
Gain
PD : 1 APD : 50
Scintillator : ()
Light yield (ph/MeV) | Decay time (ns)
50,000 | 3,000
Scintillator : ()
MPPC
Size
1mm×1mm
Pulse readout
Current readout
(Scintillator : )
Gain :
S8664-11
S12571-050C
Light yield (ph/MeV) | Decay time (ns)
10,000 | 25

## Slide 23: Comparison the CT images
X-ray : 120kV
Alcohol
()
Water
()
Phantom
PD
APD
MPPC
(Current readout)
MPPC
(Pulse readout)
CNR = 1.42
CNR = 4.41
CNR = 17.9
CNR = 5.32
Contrast-to-noise ratio
: photon statistic
~ radiation dose
Comparing the CNR, we estimated
MPPC device enable to reduce radiation dose ~ 1/100

## Slide 24: X-ray computed tomography
Projection data
Scintillator
MPPC
Phantom
Cu Collimator
X-ray generator
Al filter (1mm)
LSI
Ch 1
Ch 2
Ch 3
Ch 4
~
Ch 15
Ch 16
~
~
~
Ratio of incident X-ray
and transmitted X-ray

## Slide 25: Material identification
Line attenuation
coefficient (LAC)
[/g]
[g/]
[/cm]
Depending on the density of the materials,
Different materials will have the same CT value
Conventional CT,
CT value is the only parameter.
↓
It is difficult to discriminate materials
If the photon energy is constant,
MLAC becomes a substance-specific value
Mass line attenuation coefficient ()
Photon energy (keV)
water
iodine
Mass line attenuation
coefficient (MLAC)
Density
=
×

## Slide 26: Count rate of our system
Motivation
Measurement data
The radiation dose of CT is
It is necessary to investigate X-ray tolerance
Measured output rate 5MHz/
X-ray count rate (Tube voltage : 150kV)
Counted ~1/100 of conventional CT rate
We attempted CT imaging
under low dose rate (several MHz/)

## Slide 27: Energy resolution of our PC-CT system
spectrum ( 16ch )
Motivation
Each energy window width
is determined by energy resolution
Measurement method
-ray spectrum at 16ch
Each energy window width is 20keV
Energy discrimination is possible
Energy resolution
31%(FWHM)
Photoelectric peak @ 60keV

## Slide 28: Beam hardening artifact
CT image
Al
X-ray spectrum
Incident
X-ray
Radiation quality
① < ② < ③
Al
transmitted
X-ray
X-ray

## Slide 29: Beam hardening artifact
All energy : 30-150keV
Phantom
X-ray : 150kV

## Slide 30: Line attenuation coefficient
water
alcohol
High
Low
Contrast ratio of water and alcohol
Line attenuation coefficient [/cm]

## Slide 31: Measured CT value vs Theoretical CT value
Plastic
Iron
Aluminum
CT value [HU]
CT value [HU]
Photon energy [keV]
Photon energy [keV]
Photon energy [keV]
keV
Intensity
We measured 4 CT values of each material
from 4 energy resolved images
Choosing a proper energy window, our PC-CT can identify lighter materials
CT value [HU]
Plot : Measured CT value
Line : Theoretical CT value

## Slide 32: Color imaging of lighter phantom
20-30 keV
30-60 keV
60-90 keV
90-150 keV
(A)
(B)
(C)
(D)
(E) = (B)/(A)
(F) = (D)/(B)
To emphasize plastic parts, we compared 4 energy resolved images

## Slide 33: Color imaging of lighter phantom
20-30 keV
30-60 keV
60-90 keV
90-150 keV
(A)
(B)
(C)
(D)
(E) = (B)/(A)
(F) = (D)/(B)
To emphasize aluminum parts, we compared 4 energy resolved images
Notes: この図見せられても、伝わらない気が。。。

## Slide 34: Variation of iodine density value between 16ch
Iodine density image
5 mg/ml
10 mg/ml
15 mg/ml
20 mg/ml
…
ROI in each iodine phantom
Histogram of density value between 16-ch

## Slide 35: Variation of iodine density value between 16ch
Iodine density image
5 mg/ml
10 mg/ml
15 mg/ml
20 mg/ml
…
ROI in each iodine phantom

## Slide 36: Gadolinium K-edge imaging
Water
Gd
5mg/ml
10mg/ml
15mg/ml
20mg/ml
θ
Gd K-absorption edge : 50.2 keV
Eth 0,1,2,3 : 35, 45, 55, 70 keV
Gd concentration
5, 10, 15, 20 mg/ml
Projection of the density image
16-ch Gd density image
Using the K-absorption edge,
our PC-CT enable Gd K-edge imaging
16
1
2
3
4
1
2
3
4
20 mg/ml
15 mg/ml
10 mg/ml
5 mg/ml
Gd phantom