## Slide 1
S. K. Sundaram Multifunctional Materials Laboratory (M&M-Lab)
Kazuo Inamori School of Engineering The New York State College of Ceramics Alfred University, Alfred, NY 14802
M. Diwan
Brookhaven National Laboratory
Upton, NY 11973
Contributors: AU - Kameron Chambliss and Ruhil Dongol
BNL - Athena Marneris, Ken Sexton, and Rahul Sharma
The Sundaram Group
Photomultiplier Tube Glass Failure in Neutrino Detection: Challenges and Opportunities
Multifunctional Materials Laboratory
(M&M-Lab)

## Slide 2
Outline
2
Introduction
  Major challenges
  Our objectives
  Our approach
PMT Glasses
  Composition
  Preparation
Selected Results
  Vickers hardness
  Ring-on-ring biaxial flexural strength test
  Crack velocity measurement
Summary and Conclusion
Future Plan and Opportunities
Acknowledgment

## Slide 3
Major Challenges
3
  Photomultiplier tube (PMT) glasses will be submerged at a water depth of ~80 meters under pressure up to 880 MPa at 14oC, exposed to high purity water corrosion for about 20 years.
  New PMT bulb design is considered for the Long Baseline Neutrino Experiment Collaboration (LBNE).
  There is limited glass strength/fracture mechanics data on PMT glasses (particularly data for service under the conditions for expected LBNE campaign period) and also on chemical durability of PMT glasses
  There is no correlation between chemical durability and fracture mechanics developed for the PMT glasses.
  Full prototypic testing is cost prohibitive.
  We need to collect basic datasets to perform required statistical analysis to predict the time to failure accurately.

## Slide 4
Our Objectives
4
To conduct comprehensive study of fracture mechanics of selected PMT glasses
To predict lifetime of PMT glass under the neutrino detection conditions
To investigate new glass compositions with superior properties for better detection

## Slide 5
Our Approach
5
We used vendor glasses. We combined laboratory measurements and prototypic testing to evaluate these glasses.
Laboratory measurements (Alfred University)
    Basic glass properties
    Flexural strength
    Vickers hardness testing
    Ring-on-ring equibiaxial flexural strength testing
    Chemical durability of the glasses
    Proof-testing
Prototypic testing of PMTs (Brookhaven National Laboratory)
    Crack velocity measurements

## Slide 6
PMT - Composition
6
The nominal compositions of two vendor glasses fall in the following ranges: Sodium borosilicate glass composition
Na2O gives flexibility in melting temperature
ZnO reduces the “misty” and “milky” formation in PMT glass*
Vendor glasses: #1 ETL and #2 Hamamatsu
Oxide | SiO2 | Na2O | Al2O3 | ZnO | BaO | CaO | B2O3
Weight Percent | 65-70 | 6-9 | 4-7 | 0-3 | 0-3 | 0-1 | 15-18
* Corrosion in Glass Windows in DIRC PMTs – P. Bourgeois and J. Va’vra

## Slide 7
PMT Glass Sample Preparation
7
Melt at 1450°C for 12 hours in a platinum crucible
Anneal at 550°C for 1 hour
Grinding Process
  Average Grain size in µm 45.2, 30.4, 20.9, 14.0, 9.2, 6.2
  3-5 min on either side
Polishing Process
  Diamond suspension polycrystalline
  3µm and 0.25µm

## Slide 8
Results – Vickers Indentation
8
200g Indent
300g Indent
500g Indent
1000g Indent

## Slide 9
Results – Vickers Indentation
9
Chart: Load/g
| 200.0 | 300.0 | 500.0 | 1000.0
Vendor1 | 542.0241 | 525.5929 | 518.1029 | 507.7312
Vendor2 | 583.9164 | 553.3328 | 538.2889 | 524.5075
Vendor glasses: #1 ETL and #2 Hamamatsu
Schott Supermax® 33 Rolled Borosilicate Glass has Vickers Hardness of 568.
http://www.us.schott.com/supremax/english/product_properties/index.html
Chart: Load/g
| 200.0 | 300.0 | 500.0 | 1000.0
Vendor1 | 542.0241 | 525.5929 | 518.1029 | 507.7312
Vendor2 | 583.9164 | 553.3328 | 538.2889 | 524.5075
Vendor 2 (5days) | 563.7563 | 532.9444 | 520.1719 | 511.9741
Vendor1(5 days) | 543.2381 | 528.6969 | 515.8813 | 505.146

## Slide 10
Results – Ring-on-Ring Test
10
Equibiaxial strength, σf , of a circular plate in units of MPa is
Where F is breaking load in N

## Slide 11
Results – Ring-on-Ring Test
11
Vendor2 Sample | Diameter (mm) | Thickness (mm) | Breaking Load (N) | Equibiaxial Strength (MPa)
1 | 29.77 | 2.368 | 738.14 | 84.1
2 | 29.725 | 2.763 | 1156.65 | 96.8
3 | 29.755 | 2.473 | 702.38 | 73.4
4 | 29.82 | 2.719 | 1105.21 | 95.5
5 | 29.675 | 2.793 | 1094.73 | 89.8
Vendor1 Sample | Diameter (mm) | Thickness (mm) | Breaking Load (N) | Equibiaxial Strength (MPa)
1 | 29.60 | 3.192 | 1415.81 | 89.5
2 | 29.65 | 3.015 | 1045.38 | 74.0
3 | 29.82 | 2.735 | 1073.84 | 92.2
4 | 29.65 | 3.117 | 1378.42 | 91.3
Vendor glasses: #1 ETL and #2 Hamamatsu
Equibiaxial Strength for borosilicate glass range is 46 – 1350 MPa.
*http://www.dtic.mil/dtic/tr/fulltext/u2/a546252.pdf

## Slide 12
Results – Implosion Tests
12
Several implosion tests of PMTs of different diameters from the vendors were completed at BNL. These test results (summarized in the next slide) were used to estimate crack velocity.
At AU, the slow motion videos were reviewed frame-by-frame and visible cracks were measured in each frame. These distances were scaled to the known diameter of the bulb and the frame-rate of the video was used to calculate a velocity for each crack.
Selected images of a couple of implosion tests are also shown in the following slide. Cracks that are evolving and being measured in the current frame are indicated with red arrows. Cracks that have been measured are marked in blue.

## Slide 13
Results – Summary of Implosion Tests
13
5 minutes rise
2 hours rise
24 hours rise
Hamamatsu 10” PMT Drawing

## Slide 14
Results – Crack Velocity
14
Frame 6
5000 fps
RA0227
Crack #4
Frame 10
4800 fps
RA0246
Frame 10
6000 fps
TA3281
Frame 5
6000 fps
TA3610

## Slide 15
Results – Crack Velocity
15
In all, thirteen individual cracks were measured in the four different bulbs and an average velocity of 184.6 m/sec was calculated with a standard deviation of 58.7 m/sec
Bulb Designation | Pressure at Break (psi) | # of Cracks Measured | Average Crack Velocity (m/sec) | Standard Deviation (m/sec)
RA0246 | 260 | 4 | 191.91 | 63.70
RA0227 | 269 | 1 | 192.26 | --
TA3610 | 205 | 3 | 239.20 | 57.08
TA3281 | 103 | 5 | 144.35 | 36.52
Our data indicate large variations in the mechanical properties due to differences in flaw populations in the PMT bulbs, underlining the need for more systematic fracture mechanics testing.

## Slide 16
Summary and Conclusion
16
Vickers hardness numbers for both vendor glasses are comparable to known borosilicate glass hardness and show a non-linear decreasing trend as expected. Hardness values for Vendor 2 glass decreased considerably for sample treated in high purity water exposure test for 5 days
  Hydrolytic process may be responsible for the dramatic decrease in hardness number for Vendor 2.
Equibiaxial Strength results show comparable values to literature values of borosilicate glass hardness
Crack velocities measured are much higher than the values reported for common glasses.
Extensive testing is required to understand the fracture of these glasses and predict their failure.

## Slide 17
Future Plan and Opportunities
17
We seek to develop and understand a more complete picture of the flexural strength of these glasses, via a full series of static fatigue measurements in dry air and high-purity water.
We plan to use vapor hydration testing (designed originally to evaluate nuclear waste glasses) for accelerated testing of PMT glasses for chemical durability.
We will use proof testing to determine the time for failure data at required confidence level. Selected segments of the PMT bulbs will be used for proof-testing and statistical analysis.
Our results will advance the detection science:
  A predictive relationship that is useful in discovery of new materials
  Improved glass composition for advanced neutrino detection.

## Slide 18
Acknowledgement
18
Thanks for your time – I welcome your questions and comments
Kyocera Corporation for the support of Inamori Professorship (SKS)
Dean Doreen Edwards, Kazuo Inamori School of Engineering for support
Prof. Kathleen Richardson, Clemson University for part of travel support
Brookhaven National Laboratory for project support